## Slide 1
Clinical Training Presentation
This clinical training PowerPoint does not replace the Operator’s Manual.

## Slide 2: Learning Objectives
System Overview
Patient Assessment & Consultation
Clinical Use of the Device
Treatment Techniques

## Slide 3
System
Overview

## Slide 4: Indications for Use
The eCO2™ 3D laser system produces a 10,600 nm wavelength.
The system is indicated for use in dermatological procedures requiring ablation (removal), resurfacing, and coagulation of soft tissue. Additionally, the 120- and 300-micron spot sizes are used in the treatment of wrinkles, rhytides, furrows, fine lines, textural irregularities, pigmented lesions, and vascular dyschromia.
The 500-micron tip and the OPTIONAL F50, F100, and Zoom non-fractional handpieces are also indicated for use in skin resurfacing and surgical applications requiring the ablation, vaporization, excision, incision, and coagulation of soft tissue in medical specialties including dermatology, plastic surgery, and podiatry. See the Operator’s Manual for a full list of procedures.

## Slide 5: Coefficient Scale
The eCO2™ 3D laser system targets water using the 10,600 nm wavelength.

## Slide 6: Device Specifications
HANDPIECES & TIPS | LASER MEDIUM | AIMING BEAM | POWER (W) | PULSE RATE (Hz)
Scanner: - 120 µm spot size - 300 µm spot size - 500 µm spot size Surgical: F50 (0.1 mm)* F100 (0.2 mm)* Zoom (0.2 – 1.0 mm)* | CO2 | 520 nm** | Scanner: 5 - 40 Surgical: 0.05 - 30 | Scanner: 10 – 200 Surgical: 5 - 400
**Aiming beam is less than 5 mW. Per ANSI Class IIIA, does not require wavelength-specific eye protection. Avoid intra-beam viewing.
*Optional handpieces

## Slide 7
Patient
Consultation &
Assessment

## Slide 8
Fitzpatrick Skin Type

## Slide 9: Determine Suitability
In determining suitability, providers should consider the following factors for each individual patient:
Fitzpatrick skin type
Sun exposure
Age
Overall skin health
Medical history
Current medications
Reason for seeking treatment
Expectations

## Slide 10: Contraindications
NOTE: The scanner handpiece tip contains a magnet. Patients with a pacemaker, an implantable defibrillator, or other electronic medical device should not approach or receive treatment with the scanner handpiece. Electromagnetic forces can interfere with these devices.
Pregnancy
Open wound, inflammation, or infection in the area to be treated
Cancer or other malignant diseases
Systemic disease or condition
A bleeding disorder
Photosensitive skin or taking medications that cause photosensitivity
Skin lesions caused by diabetes mellitus, connective tissue disorders, radiation therapy, or chemotherapy
Skin allergy to local anesthetics, antibiotic medication, or any other types of pertinent medications

## Slide 11: Contraindications
History of keloid scarring or connective tissue disease
Active breakout of herpes simplex
Use of immunosuppressive drugs, any other medications, and/or diseases which can affect the wound-healing process
Recent Vitamin A derivatives use
Psycho-neurotic patients, including alcoholics or drug abusers
Patients who are unable or unwilling to follow the post-treatment instructions
Patients with unrealistic expectations regarding the clinical outcome of the treatment

## Slide 12: Pre-Treatment Guidance
Review the cautions and contraindications listed above and advise the patient accordingly.
To reduce the risk of hyperpigmentation, consider using a product to suppress melanocyte activity (e.g., hydroquinone) and daily sunscreen for two weeks before and after the treatment, especially in higher Fitzpatrick Skin Types.
Avoid the following for at least two weeks before treatment:
  Anticoagulants (e.g., NSAIDs, Eliquis®)
  Iron supplements
  Herbal supplements (e.g., ginkgo, ginseng, garlic, St. John's Wort)
  Tetracycline antibiotics (e.g., doxycycline)
  Any medications that cause photosensitivity (e.g., amiodarone, hydrochlorothiazide)
  Sun exposure, tanning bed use, spray tanning, or tanning products
  Neuromodulator injections in the treatment area
  Any professional skin treatment in the treatment area
  Retinol, AHAs, or other active topicals
This information has been developed in collaboration with operators who are experienced in using the device and/or from medical literature. It is only intended as a general guide.

## Slide 13: Pre-Treatment Guidance
Avoid receiving dermal fillers in the treatment area at least six weeks before treatment.
Patients with a history of herpes simplex may obtain prophylactic medication from a prescriber.
Arrive for the treatment with clean skin. Do not apply skincare products, lotion, makeup, scents, or sunscreen to the treatment area.
If the treatment includes a hair-bearing area (e.g., beard), arrive for the treatment freshly shaved.
This information has been developed in collaboration with operators who are experienced in using the device and/or from medical literature. It is only intended as a general guide.

## Slide 14: Post-Treatment Guidance
Patients may experience a burning sensation for 1 to 2 hours immediately after the treatment. Cooling with an ice pack can provide relief to the treated area. Enclose the ice pack in a sterile and dry gauze to prevent water from permeating onto the treated area. Do not apply ice pack directly to the skin.
Do not let the treatment area be in contact with water for up to 24 hours after the procedure. After 24 hours, the patient may gently wash and dry the treatment area using a patting motion. Do not rub the area. Do not scrub or pick the skin.
Some patients might experience pinpoint bleeding. If this occurs, gently pat the area using a clean gauze to stop bleeding.
Keep the skin moisturized and hydrated after the treatment. Apply an ointment or barrier cream product to the treated area several times a day for the next 3 to 5 day or until re-epithelialization is complete. This will promote healing and reduce adverse reactions.
After 36 hours, patients may apply mineral makeup, shave, and wash regularly.
This information has been developed in collaboration with operators who are experienced in using the device and/or from medical literature. It is only intended as a general guide.

## Slide 15: Post-Treatment Guidance
A thin crust may appear within 3 to 7 days post-treatment as new skin begins to form. The new skin is sensitive and must be protected from sunlight. Sunscreen should be applied generously on the treated area.
Avoid using alcohol-containing or harsh cosmetics for at least 2 weeks after the treatment.
Avoid vigorous activity for the first week post-treatment or until initial healing has occurred. Increased redness and irritation may result from any activity that increases blood flow or body temperature (i.e., alcohol consumption, exercise, or sauna).
Over the next month, keep applying moisturizer and sunscreen with SPF 30+ and use an umbrella, cap, or any available protections against sunlight.
Patient should follow all post-treatment care outlined by the treatment provider.
This information has been developed in collaboration with operators who are experienced in using the device and/or from medical literature. It is only intended as a general guide.

## Slide 16
Clinical Use

## Slide 17: Reference Materials
Always refer to the Operator’s Manual for information regarding the product.
Cynosure Lutronic assumes no liability whatsoever for any damage, loss, or injury resulting from an application of a product which is not in strict accordance with its Operator’s Manual. Cynosure Lutronic also assumes no liability for any damage or injury arising as a result of operator error or mistake.
Usage of the product is based on the practitioner’s knowledge and experience. The practitioner is solely responsible for treatment results. There is no guarantee that the expected or anticipated results will be achieved.
Training conducted on the device and all treatments performed should adhere to the guidelines found within the Operator’s Manual.

## Slide 18: Electromagnetic Spectrum
The eCO2™ 3D laser system targets water using the 10,600 nm wavelength.

## Slide 19
Treatment
Techniques

## Slide 20
Scanner
Handpiece

## Slide 21: Scanner Handpiece Operation Modes
STATIC mode:
  Laser will fire a single complete pulse if either the footswitch is depressed, or the Skin Sensor is enabled, and the tip is in contact with the skin while the footswitch is depressed.
  Keep the handpiece stationary until each laser pulse is complete.
  Used for the stamping technique.
DYNAMIC mode:
  Laser will fire continuously while either the footswitch is depressed, or the Skin Sensor is enabled, and the tip is in contact with the skin while the footswitch is depressed.
  Keep the handpiece perpendicular and adjust the angle of your hand to follow the contour of the patient’s skin as you move across the treatment area.
  Used for the painting/airbrush technique.
This information has been developed in collaboration with operators who are experienced in using the device and/or from medical literature. It is only intended as a general guide.
Correct Placement
Incorrect Placement

## Slide 22: Scanner Handpiece Tips
The 120 µm tip generates deep and narrow Micro Ablative Columns (MACs), ideal for scars, deep wrinkles, and tightening.
The 300 µm tip generates mid-depth and size MACs, ideal for pigment and fine texture.
The 500 µm tip generates shallow and wide MACs, ideal for skin tone, fine texture, and luminosity.
This information has been developed in collaboration with operators who are experienced in using the device and/or from medical literature. It is only intended as a general guide. Images not to scale. For illustrative purposes only.

## Slide 23: Treatment Guidelines
Determine the patient's Fitzpatrick Skin Type and treatment goals. Using this information, reference the corresponding settings in the Operator's Manual and/or Quick Start Guide.
Perform test spots at several energy levels to ensure that the expected response is observed. When the patient returns for full treatment several days up to two weeks later, observe that the healing process is as expected. Adjust the treatment parameters accordingly.
If using multiple tips in the same treatment area, use the tip with the deepest depth first.
Higher peak power is suggested for acne scars due to the corresponding short pulse duration to ablate the tissue deeply.
Lower peak power is suggested for wrinkles due to the corresponding long pulse duration, generating a more significant thermal effect.
Lower density and energy settings are suggested for the periorbital, neck, and chest areas.
Do not treat over the patient's eyebrows, hairline, or beard hair.
This information has been developed in collaboration with operators who are experienced in using the device and/or from medical literature. It is only intended as a general guide.

## Slide 24: Treatment Guidelines
If primarily treating with Static mode, follow the treatment with Dynamic mode around the edges to blend the result with the nearby untreated skin.
Treatment Plan:
  1 to 3 sessions for anti-aging
  3 to 5 sessions for scars
  4 to 6 weeks between sessions
  Regular follow-up visits
Maximal results are generally observed 4 to 6 months after the final treatment.
This information has been developed in collaboration with operators who are experienced in using the device and/or from medical literature. It is only intended as a general guide.
Image not to scale. For Illustrative purposes only.

## Slide 25: Precision Peel Treatment
Safe and effective for FST I to IV, the eCO Tone™ protocol is used to perform the Precision Peel, which lightly resurfaces the skin to improve fine lines and sun damage with the 120, 300, or 500 µm tip.
This treatment maximizes comfort while minimizing impact on the patient's daily routine and lifestyle.
Patients typically experience only 2 to 3 days of mild redness.
A topical anesthetic is not generally required. Pair with the Zimmer® chiller for enhanced patient comfort.
This information has been developed in collaboration with operators who are experienced in using the device and/or from medical literature. It is only intended as a general guide.

## Slide 26: Before & After

## Slide 30: Thank You
Intended for distribution in US only.
Cynosure, LLC and Lutronic Aesthetic, Inc. do not diagnose or treat medical conditions, prescribe treatment protocols, provide medical advice, or otherwise practice medicine. Patient should consult with a qualified healthcare professional for diagnosis and proper treatment protocols/parameters to determine if treatment with any product is right for them. Qualified healthcare professionals are responsible for diagnosis and treatment protocols using products.
©2025 CYNOSURE LUTRONIC, INC. All Rights Reserved. eCO2™ 3D and Lutronic is a registered trademark of Lutronic Corporation. Cynosure is a registered trademark of Cynosure, LLC.